नीचे ViscoLogic Viscosity Meter का Complete, Updated, Expandable Software Architecture है, जैसा हमने discuss किया। इसमें कुछ भी missing नहीं रखा: Auto-open + power-fail recovery, one-time commissioning password, Tabletop/Inline modes, Local/Remote/Mixed control (PLC + UI), Modbus TCP data out + commands in, 150 mA safety, reading health % (No ML), unlimited calibration points, UI + backend + logging + retention, और future expandability।

# 1) System का simple purpose

Software का काम:

- Drive coil को safe तरीके से चलाना

- Pickup coil signal ADC से पढ़ना

- Noise हटाकर stable measurement निकालना

- Auto-tune करके resonance पर lock रहना

- Temperature read करके correction करना

- Calibration model से viscosity (cP) निकालना

- UI पर दिखाना और data log करना

- PLC को values भेजना और PLC से start/stop/mode commands लेना (अगर remote enabled)

# 2) High-Level Layered Architecture (समझने लायक)

UI Layer (PyQt)

- Operator Screen

- Engineer Screen (password)

- Calibration Wizard

- Alarms + Diagnostics

|

Core Orchestrator (State Machine + Rules)

- Startup/Self-check

- Commissioning gate

- Mode + Control arbitration

- Safety manager

- Measurement scheduling

|

Measurement Pipeline (DSP + Model)

- Sweep + Lock (PLL-like)

- Digital Lock-in (I/Q)

- Health score (No ML)

- Viscosity compute + Temp compensation

|

Drivers (Hardware Abstraction)

- ADS1115 differential + PGA auto-range

- PWM drive + soft-start

- PT100/MAX31865

- Optional: current/thermal monitor

|

Protocol Adapter

- Modbus TCP Server (PLC reads/writes)

|

Storage/Logging

- SQLite (profiles, state, events)

- CSV logs + retention cleanup

# 3) Boot + Auto-Open + Power Fail Recovery (Must)

## 3.1 Auto start on boot

- Raspberry Pi boot होते ही app auto-run (systemd service)

- crash होने पर auto-restart

- startup पर always SAFE IDLE (drive OFF) जब तक checks pass न हों

## 3.2 Power fail recovery

- Power आते ही app फिर auto start

- self-check

- last settings restore (mode/control source/log enable)

- Tabletop में default: user start करेगा

- Inline में (अगर enabled): self-check OK के बाद auto-run resume

# 4) One-time Commissioning Password (New machine only)

यह feature “नई machine” पर सिर्फ पहली बार password मांगेगा।

## 4.1 Commissioning flow

- Boot → Self-check

- अगर commissioned=false:

  - Screen: Commissioning Lock (password)

  - सही password के बाद: Commissioning Wizard

  - Wizard complete → commissioned=true

- अगली बार से commissioning password नहीं पूछेगा

## 4.2 Engineer password अलग रहेगा (recommended)

- Engineer screen (settings/calibration) हमेशा password protected

- Commissioning password एक बार का है, engineer password regular protection है

## 4.3 Reset / Re-commission

Engineer-only option:

- “Re-commission device” (commissioned=false)

- फिर next boot पर commissioning password फिर मांगेगा

# 5) Modes: Tabletop vs Inline (Online)

## 5.1 Tabletop Mode (Lab/sample)

- Operator dip करता है, Start/Stop manual

- Stop = measurement बंद

- Calibration friendly

## 5.2 Inline Mode (Plant continuous)

- 24x7 running behavior

- Control typically PLC से

- UI में “Enable / Pause (Maintenance)”

- Auto relock + alarm handling continuous

- Comm-loss behavior configurable (Pause recommended)

UI में toggle:

- Mode: Tabletop / Inline

# 6) Control Source: Local / Remote / Mixed (सबसे जरूरी logic)

यह ensure करता है कि PLC और local UI में conflict न हो।

## 6.1 Local

- Start/Stop केवल UI से

- PLC commands ignore

- PLC values read करता है

## 6.2 Remote

- Start/Stop PLC से

- UI start disabled (view + maintenance + E-Stop allowed)

- PLC process driven

## 6.3 Mixed

- Start UI या PLC दोनों से

- Stop UI या PLC दोनों से

## 6.4 Priority rules (Final locked)

- Emergency Stop highest priority

- STOP हमेशा START को override करेगा

- Remote mode में local start disabled

- Comm loss (Inline):

  - Option A: Keep running + alarm

  - Option B (recommended): Pause (Safe) + alarm

- Fault latch states में start blocked until reset

# 7) Safety Architecture (150 mA safe operation)

आपका point सही था: device safe रखना mandatory है।

## 7.1 Hardware safety (Must)

- Drive stage में current limit (hardware enforced)

- Absolute max: 150 mA

- Fuse/PTC optional, thermal consideration

## 7.2 Software safety (Must)

- Startup: Drive OFF until READY

- Soft-start ramp (0 → target smoothly)

- Air calibration = low power:

  - 30–60 mA

  - max 10–15 sec

  - auto stop

- Overcurrent/overheat/saturation पर:

  - Drive OFF

  - Alarm + latch (as configured)

- ADC saturation detection:

  - PGA auto adjust + warning

- Inline comm loss action enforced

# 8) Measurement Engine (Stable reading, best practical)

## 8.1 Acquisition (ADS1115 differential)

- A0-A1 differential read

- PGA auto-range to avoid clipping

- Sampling within ADS1115 limits, output stable via DSP averaging

## 8.2 Auto sweep + resonance lock (PLL-like)

- Startup sweep: 150–200 Hz (configurable)

- Best region find, then fine tune

- Runtime drift tracking:

  - lock breaks तो auto re-sweep

- Lock states:

  - SEARCHING → LOCKING → LOCKED → LOST_LOCK (auto retry)

## 8.3 Digital lock-in (I/Q) for noise rejection

- Reference sin/cos at drive frequency

- I/Q demod + low-pass

- magnitude + phase extracted (stable)

## 8.4 Reading Health % (No ML)

यह ML नहीं है, deterministic quality score है:

- lock stability

- signal strength vs noise floor

- clipping/near clipping

- variance (last N seconds)

- lost lock count
Output: 0–100%

Operator-friendly label:

- “Reading Health %” या “Signal OK %”

# 9) Calibration System (Unlimited points, multiple profiles)

## 9.1 Why

हर plant/fluid अलग होता है, इसलिए calibration flexible होना चाहिए।

## 9.2 Wizard flow

- Add calibration point:

  - sample name

  - known viscosity (cP)

  - temperature auto record

  - capture 10 sec stable

  - save point

- Unlimited points add

- Multiple profiles support (Oil-A, Oil-B, Customer-1)

## 9.3 Model (Safe default)

- Default: LUT + interpolation

- Advanced options future: polynomial/other fits (engineer mode only)

# 10) Temperature Handling

- PT100 via MAX31865

- temp smoothing

- output:

  - As-measured viscosity

  - Optional normalized viscosity (profile-based)

# 11) UI Specification (Everyone can use)

## 11.1 Operator Screen (Simple)

- Big: Viscosity (cP)

- Temp (°C)

- Status light: Searching/Locked/Error

- Reading Health %

- Buttons (mode dependent):

  - Tabletop: Start / Stop

  - Inline: Enable / Pause (Maintenance)

- Logging: Start Log / Stop Log

- Export

- Alarm Details

## 11.2 Commissioning Screens (One time)

- Commissioning Lock (password)

- Commissioning Wizard:

  - Mode select

  - Control source select

  - Remote enable

  - comm-loss action

  - safety limits view

  - complete setup

## 11.3 Engineer Screen (Password)

Tabs:

- Mode & Control

- Calibration Wizard + profile manager

- Diagnostics (ADC, temp, saturation, lock stats)

- Trends (viscosity/freq/health)

- PLC/Network (Modbus settings + register map view/export)

- Maintenance tools (re-commission, reset alarms, log retention)

# 12) PLC Integration (Modbus TCP) Data out + Commands in

## 12.1 Recommended model

- Pi = Modbus TCP Server

- PLC = Modbus Client

## 12.2 Scaling rule

- Viscosity_x100 (int32) so decimals preserved

- Temp_x100, Freq_x100 (int16)

- Health 0–100

## 12.3 Register Map (Human address style)

PLC engineer 0-based/1-based adjustment करेगा, structure same रहेगा।

### PLC Reads (Meter → PLC) Holding Registers

| Addr | Name | Type | Example |
| --- | --- | --- | --- |
| 40001-40002 | VISC_CP_x100 | int32 | 45.23 cP → 4523 |
| 40003 | TEMP_C_x100 | int16 | 42.15°C → 4215 |
| 40004 | FREQ_HZ_x100 | int16 | 178.95 Hz → 17895 |
| 40005 | READING_HEALTH | uint16 | 0–100 |
| 40006 | STATUS_WORD | uint16 | bit-coded |
| 40007 | ALARM_WORD | uint16 | bit-coded |
| 40008 | HEARTBEAT | uint16 | every 1 sec +1 |
| 40009 | LAST_CMD_SEQ | uint16 | last processed |

### PLC Writes (PLC → Meter) Holding Registers

| Addr | Name | Type | Meaning |
| --- | --- | --- | --- |
| 41001 | MODE_SELECT | uint16 | 0 Tabletop, 1 Inline |
| 41002 | CONTROL_SOURCE | uint16 | 0 Local, 1 Remote, 2 Mixed |
| 41003 | REMOTE_ENABLE | uint16 | 0/1 |
| 41004 | CMD_WORD | uint16 | bits below |
| 41005 | CMD_SEQ | uint16 | increment each new cmd |
| 41006 | CMD_RESULT | uint16 | 0 OK else error code |
| 41007 | LOG_ENABLE | uint16 | 0/1 |
| 41008 | PROFILE_ID | uint16 | select profile |

### CMD_WORD bits (PLC → Meter)

- bit0 START_REQUEST

- bit1 STOP_REQUEST

- bit2 PAUSE_MAINTENANCE

- bit3 RESET_ALARMS

- bit4 CLEAR_ESTOP_LATCH (अगर safety allows)

### STATUS_WORD bits (Meter → PLC)

- bit0 READY

- bit1 SELF_CHECK_FAIL

- bit2 SEARCHING

- bit3 LOCKED

- bit4 LOST_LOCK

- bit5 RUNNING (drive active)

- bit6 PAUSED/MAINTENANCE

- bit7 REMOTE_ACTIVE

- bit8 ALARM_ACTIVE

- bit9 ESTOP_LATCHED

- bit10 ADC_SATURATION

- bit11 TEMP_SENSOR_FAULT

- bit12 COMMISSIONING_REQUIRED

### ALARM_WORD bits (example)

- bit0 ADC_NOT_FOUND

- bit1 TEMP_SENSOR_FAIL

- bit2 PICKUP_SIGNAL_TOO_LOW

- bit3 SIGNAL_CLIPPED / PGA_LIMIT

- bit4 DRIVE_OVERCURRENT

- bit5 DRIVER_OVERHEAT

- bit6 LOCK_TIMEOUT

- bit7 PLC_COMM_LOSS

## 12.4 Handshake (no confusion)

- PLC writes CMD_WORD + CMD_SEQ

- Meter executes once per new CMD_SEQ

- Meter updates LAST_CMD_SEQ + CMD_RESULT

# 13) Storage, Logging, Retention (Backend)

## 13.1 SQLite tables (recommended)

- device_state (commissioned flag, commissioned_at)

- settings (mode/control source, limits, comm-loss behavior)

- calibration_profiles

- calibration_points

- events (alarms, state changes)

- optional measurement history (if needed)

## 13.2 CSV logging

- timestamp, viscosity, temp, freq, health, status, alarms

- retention cleanup:

  - keep last N days

  - auto delete old logs

# 14) Expandable Design (Future ready)

### 14.1 Hardware interfaces (easy replacement)

- IAdcReader (ADS1115 today, faster ADC tomorrow)

- IDriveController (PWM today, DAC tomorrow)

- ITempSensor (MAX31865 today, other tomorrow)

### 14.2 Protocol adapters

- Modbus TCP today

- OPC-UA future

- MQTT future

- REST API optional

### 14.3 Multi-sensor support future

- same orchestrator, multiple instances

- per-sensor register blocks in PLC mapping

# 15) Suggested Folder Structure (Implementation ready)

viscologic/

app.py

core/

orchestrator.py

state_machine.py

safety_manager.py

diagnostics.py

config_manager.py

event_bus.py

security/

commissioning_manager.py

auth_engineer.py

drivers/

adc_ads1115.py

drive_pwm.py

temp_max31865.py

dsp/

lockin_iq.py

sweep_tracker.py

filters.py

health_score.py

model/

calibration_store.py

calibration_lut.py

viscosity_compute.py

temp_compensation.py

protocols/

modbus_server.py

register_map.py

storage/

sqlite_store.py

csv_logger.py

retention.py

ui/

operator_screen.py

engineer_screen.py

calibration_wizard.py

alarms_screen.py

commissioning_lock.py

commissioning_wizard.py

deploy/

viscologic.service

config/

settings.yaml

schema.json

tests/

test_state_machine.py

test_register_map.py

test_lockin.py

test_sweep.py

# 16) Development Roadmap (Build order)

- systemd auto-start + safe idle + self-check

- drivers: ADS1115 + PT100 + PWM

- state machine + safety manager

- sweep + lock + stable outputs

- reading health + alarms

- calibration wizard + profiles + LUT

- operator UI + engineer UI

- Modbus TCP server + register map + handshake

- long-run testing + power fail + comm loss validation

- packaging + SOP docs

# 17) Final “Operator promise” (locked)

“Operator को सिर्फ Start/Enable करना है; system खुद safe तरीके से run होकर resonance lock करेगा, stable viscosity number दिखाएगा, data log करेगा और PLC को live values देगा।”

अगर आप चाहें तो मैं इसी architecture के आधार पर अगला step बिना सवाल के दे सकता हूँ:

- Exact PLC tag list (Siemens/AB/Delta style)

- या UI wireframe text spec (हर button/label का final wording)

- या state machine transition table (pure implementation spec)