नीचे ViscoLogic Viscosity Meter Software का complete, updated, end-to-end plan दिया है, जिसमें हमारी सारी modifications शामिल हैं: Auto-start on power, power-fail recovery, Tabletop + Inline modes, Local/Remote/Mixed control, PLC communication (Modbus TCP), safety (150 mA limit), calibration (unlimited points), confidence/reading health (No ML), logging + backend + UI। यह plan ऐसे लिखा है कि non-technical भी समझ सके और developer के लिए भी implementation ready हो।

## 1) Software का काम क्या है (Simple)

यह software:

- Drive coil को चलाता है (rod/fork vibrate करवाता है)

- Pickup coil का signal ADC से पढ़ता है

- Noise हटाकर stable measurement निकालता है

- Auto-tuning करके resonance पर lock रहता है

- Temperature पढ़कर correction करता है

- Calibration से signal को viscosity (cP) में convert करता है

- UI पर दिखाता है, data log करता है, और PLC को data देता है

- PLC से Start/Stop/Modes commands भी लेता है (अगर remote control enabled है)

## 2) Startup Behavior (Power ON, Auto Open, Power Fail)

### 2.1 Auto-Open on Power

- Raspberry Pi ON होते ही software खुद auto start होगा

- पहली screen पर Self-Check चलेगा

- Default state: SAFE IDLE (Drive OFF)

### 2.2 Power Fail Recovery

Power वापस आते ही:

- software फिर से auto open

- self-check

- last settings restore (mode, control source, logging option)

### 2.3 Auto Resume Rule (Mode wise)

- Tabletop Mode: Auto resume = OFF (operator Start करेगा)

- Inline Mode: Auto resume = ON (self-check OK होने पर auto run)
यह engineer setting से change भी हो सकेगा।

## 3) Operating Modes (Tabletop vs Inline)

### 3.1 Tabletop Mode (Lab / Sample Testing)

- operator dip करेगा, Start/Stop करेगा

- calibration आसान

- “Stop” = measurement fully stop

- PLC optional (mostly read-only)

### 3.2 Inline Mode (Online / Continuous)

- system 24x7 continuous चलेगा

- “Stop” की जगह:

  - Pause/Maintenance (drive OFF, status ON)

  - Emergency Stop (hard stop + latched alarm)

- auto relock + auto retry enabled

- PLC integration recommended

UI में साफ toggle:

- Mode: Tabletop / Inline

## 4) Control Source (Local / Remote / Mixed) — Start/Stop कौन करेगा?

यह key requirement है ताकि PLC और local UI में conflict न हो।

### 4.1 Local Control

- Start/Stop केवल meter UI से

- PLC commands ignore

- PLC केवल reading ले

### 4.2 Remote Control

- Start/Stop PLC से

- UI buttons locked/disabled (view only)

- Local पर केवल maintenance + emergency stop (optional)

### 4.3 Mixed Control

- Start UI या PLC दोनों से possible

- Stop UI या PLC दोनों से possible

- Priority rules lock होंगे (नीचे)

UI में setting:

- Control Source: Local / Remote / Mixed

- Remote Enable: ON/OFF (extra safety)

## 5) Safety Rules (Drive current up to 150 mA, device protection)

### 5.1 Hard Safety (Must)

- Drive coil current limit enforce (hardware + software)

- Absolute max: 150 mA

- Overcurrent/overheat/fault पर Drive OFF + Alarm

### 5.2 Software Safety Behaviors

- SAFE IDLE: startup पर drive OFF

- Soft-start ramp: drive धीरे-धीरे increase

- Air Calibration Low Power:

  - current low (30–60 mA)

  - time limit (10–15 sec auto stop)

- Saturation guard: ADC clip दिखे तो PGA auto adjust + warning

- Comm loss action (Inline):

  - option A: Keep running + alarm

  - option B (recommended): Pause (safe state) + alarm

## 6) Self-Check (हर startup पर)

Self-check में software verify करेगा:

- ADS1115 connected है या नहीं

- PT100 sensor OK है या नहीं

- pickup signal baseline sensible है या नहीं

- (optional) coil open/short hints

- PLC comm status (inline deployments में)

Result:

- System Ready या Self Check Failed with clear reason + Retry button

## 7) Measurement Engine (Smart part, but user-friendly)

### 7.1 Auto Tuning (Resonance Lock)

- startup पर frequency sweep (example 150–200Hz)

- best resonance point identify

- runtime में drift होने पर frequency खुद adjust
UI status:

- Searching → Locking → Locked → Lost Lock (auto retry)

### 7.2 Noise Rejection (Stable reading)

- signal cleaning algorithm चलती रहेगी ताकि hum/pump noise कम हो

- output stable बने

### 7.3 Reading Health % (Confidence, No ML)

यह ML नहीं है। यह system की reading quality बताने वाला score है:

- lock stability, signal strength, noise level, clipping, lock breaks
Operator को पता चलता है reading भरोसेमंद है या नहीं।

Label options:

- “Reading Health %” (recommended)

- “Signal OK %”

## 8) Calibration System (Unlimited points, easy process)

### 8.1 क्यों जरूरी

हर company के fluids अलग होते हैं, इसलिए calibration flexible होना चाहिए।

### 8.2 Calibration Wizard (Engineer mode)

Minimum recommended:

- Air (baseline)

- Water (1 cP)

- 2–3 standard oils (example 10, 50, 200 cP)

Best accuracy के लिए:

- user unlimited calibration points add कर सके

### 8.3 Easy workflow

- “Add Calibration Point”

- Sample name + known viscosity (cP)

- Capture 10 sec (auto stable capture)

- Save point

- Curve/LUT auto update

Model approach:

- LUT + interpolation (safe, stable, no crazy extrapolation)

- Profiles: “Oil-A”, “Oil-B” जैसी multiple profiles support

## 9) Temperature Handling

- PT100 से temperature read

- display: current temperature

- compensation:

  - basic: “as measured viscosity”

  - optional: “normalized to 25°C” (profile constants के साथ)

Engineer screen में:

- fluid profile settings (future expand)

## 10) UI Design (Easy for everyone)

### 10.1 Operator Screen

- Big number: Viscosity (cP)

- Temp (°C)

- Status light: Locked/Searching/Error

- Reading Health %

- Buttons:

  - Start/Stop (Tabletop)

  - Enable/Pause (Inline)

  - Start Log / Stop Log

  - Export

  - Alarm Details

### 10.2 Engineer Screen (Password)

- Mode select (Tabletop/Inline)

- Control source select (Local/Remote/Mixed)

- Remote enable ON/OFF

- Calibration wizard

- Settings: sweep range, averaging, limits

- Diagnostics

- Live graphs (optional but useful): viscosity trend, freq trend, health trend

## 11) Logging + Backend (Local reliable system)

### 11.1 Logging

- CSV logging (default)

- optional SQLite database (recommended for history/search)

- retention: last X days keep, old auto delete

### 11.2 Exports

- Export CSV

- optional report summary (daily/shift)

### 11.3 Crash-safe

- auto restart on crash

- last settings restore

- safe state on restart

## 12) PLC Integration (Data out + Commands in)

### 12.1 Protocol (Recommended)

- Modbus TCP

  - Pi = Modbus Server

  - PLC = Modbus Client

### 12.2 PLC reads (Outputs from meter)

- Viscosity_x100

- Temp_x100

- Freq_x100

- ReadingHealth_% (0–100)

- StatusWord (bit-coded)

- AlarmWord (bit-coded)

- Heartbeat counter

### 12.3 PLC writes (Commands to meter)

- Start_Request

- Stop_Request

- Pause/Maintenance

- Reset_Alarms

- ControlSource (Local/Remote/Mixed)

- RemoteEnable

### 12.4 Priority rules (No-conflict)

- Emergency stop highest priority

- Stop always overrides start

- Remote mode में local start disabled

- Comm loss action configurable (pause recommended)

## 13) “How operator uses it” (One page SOP)

### Tabletop SOP

- Power ON → software auto open

- (अगर needed) Air calibration

- sample dip

- Start measurement

- Locked होने पर reading नोट करो

- जरूरत हो तो log ON

- Stop

### Inline SOP

- Power ON → auto open → self-check

- Remote/Inline enabled हो तो system auto run

- PLC continuous reading ले

- Maintenance में Pause

- Alarm आए तो operator alarm details देखे

## 14) Development Execution Roadmap (Build order)

- Auto-start service + self-check + safe idle

- ADS1115 read + PGA auto-range + PT100 read

- Auto sweep + lock + stable outputs

- Reading Health % + alarms

- Calibration wizard + profiles + LUT

- Operator UI + Engineer UI

- Logging + export + retention

- Modbus TCP (PLC in/out) + priority rules

- Long-run testing + fail recovery validation

## 15) Final Promise (Locked)

“Operator को सिर्फ Start/Enable करना है; software खुद signal clean करेगा, resonance lock करेगा, stable viscosity number दिखाएगा, data save करेगा, और PLC को live values देगा।”