## ViscoLogic Viscosity Meter — Single-Page System Spec (UI + PLC + Rules)

### 1) Software का core behavior (सबको समझ आए ऐसा)

- Power ON होते ही auto-open (Raspberry Pi boot पर app खुद चालू)

- पहले Self-Check (ADC/PT100/basic signal)

- Default state: SAFE IDLE (Drive OFF)

- Tabletop mode: user “Start” दबाएगा

- Inline mode: अगर “Auto-Resume/Remote-Enable ON” है तो self-check OK होते ही auto run

- Power fail के बाद भी यही flow: auto-open → self-check → safe → (mode के हिसाब से resume)

### 2) Modes (Tabletop vs Inline)

Tabletop (Manual):

- Start/Stop manual

- Calibration आसान

- Stop = measurement बंद

Inline (Continuous):

- 24x7 running (PLC/remote control recommended)

- UI पर “Enable / Pause (Maintenance)”

- Emergency stop अलग (latched)

### 3) Control Source (Local / Remote / Mixed)

Local: Start/Stop सिर्फ meter UI से, PLC command ignore
Remote: Start/Stop PLC से, UI start disabled (view + maintenance + estop)
Mixed: Start UI/PLC दोनों से, Stop UI/PLC दोनों से (नीचे rules लागू)

Priority Rules (Conflict avoid):

- Emergency Stop सबसे ऊपर (Drive OFF + latched)

- Stop हमेशा Start को override करेगा

- Remote mode में local start disabled

- PLC comm loss (Inline) पर configurable:

  - Option A: Keep running + alarm

  - Option B (recommended): Pause (safe) + alarm

### 4) Safety (150 mA protection)

- Absolute max drive current: 150 mA (hardware + software limit)

- Air Calibration हमेशा Low-Power: 30–60 mA, max 10–15 sec, auto stop

- Soft-start ramp (sudden jerk नहीं)

- Overcurrent/overheat/ADC saturation पर Drive OFF + Alarm

## 5) PLC Integration (Recommended: Modbus TCP)

Best default: Raspberry Pi = Modbus TCP Server, PLC = Modbus Client

- PLC reading लेगा (viscosity/temp/status)

- PLC commands दे सकेगा (start/stop/mode/enable)

Note: Modbus addressing PLC-to-PLC अलग हो सकता है (कुछ PLC 40001 को “0” offset मानते हैं)। नीचे “Human address” style दिया है; PLC engineer mapping adjust कर लेगा।

## 6) Register Map (Simple + Robust)

### 6.1 PLC Reads (Outputs from meter) — Holding Registers (4xxxx)

| Human Addr | Name | Type | Scaling / Meaning |
| --- | --- | --- | --- |
| 40001–40002 | VISC_CP_x100 | int32 | 45.23 cP ⇒ 4523 |
| 40003 | TEMP_C_x100 | int16 | 42.15°C ⇒ 4215 |
| 40004 | FREQ_HZ_x100 | int16 | 178.95 Hz ⇒ 17895 |
| 40005 | READING_HEALTH | uint16 | 0–100 (%) |
| 40006 | STATUS_WORD | uint16 | bit-coded (नीचे) |
| 40007 | ALARM_WORD | uint16 | bit-coded (नीचे) |
| 40008 | HEARTBEAT | uint16 | हर 1 sec में +1 (rollover ok) |
| 40009 | LAST_CMD_SEQ | uint16 | last processed command seq |

### 6.2 PLC Writes (Commands to meter) — Holding Registers (41xxx)

| Human Addr | Name | Type | Meaning |
| --- | --- | --- | --- |
| 41001 | MODE_SELECT | uint16 | 0=Tabletop, 1=Inline |
| 41002 | CONTROL_SOURCE | uint16 | 0=Local, 1=Remote, 2=Mixed |
| 41003 | REMOTE_ENABLE | uint16 | 0=Disable remote run, 1=Enable |
| 41004 | CMD_WORD | uint16 | bit commands (नीचे) |
| 41005 | CMD_SEQ | uint16 | PLC हर नए command पर +1 करता है |
| 41006 | CMD_RESULT | uint16 (R/O) | 0=OK, else error code |
| 41007 | LOG_ENABLE | uint16 | 0=off, 1=on |
| 41008 | PROFILE_ID | uint16 | calibration profile select |

### CMD_WORD bits (PLC → Meter)

- bit0 = START_REQUEST

- bit1 = STOP_REQUEST

- bit2 = PAUSE_MAINTENANCE

- bit3 = RESET_ALARMS

- bit4 = CLEAR_ESTOP_LATCH (अगर safety allow करे)

Handshake rule:
PLC पहले CMD_WORD + CMD_SEQ लिखे → Meter process करे → 40009 (LAST_CMD_SEQ) update + 41006 (CMD_RESULT) set

## 7) STATUS_WORD (Meter → PLC) bits

- bit0 READY (self-check OK)

- bit1 SELF_CHECK_FAIL

- bit2 SEARCHING

- bit3 LOCKED

- bit4 LOST_LOCK

- bit5 RUNNING (drive active)

- bit6 PAUSED / MAINTENANCE

- bit7 REMOTE_ACTIVE (control source remote/mixed + remote enable)

- bit8 ALARM_ACTIVE

- bit9 ESTOP_LATCHED

- bit10 ADC_SATURATION (clipping/near clip)

- bit11 TEMP_SENSOR_FAULT

## 8) ALARM_WORD (Meter → PLC) bits (example set)

- bit0 ADC_NOT_FOUND

- bit1 TEMP_SENSOR_FAIL

- bit2 PICKUP_SIGNAL_TOO_LOW

- bit3 PICKUP_CLIPPED / PGA_LIMIT

- bit4 DRIVE_OVERCURRENT

- bit5 DRIVER_OVERHEAT (अगर available)

- bit6 LOCK_TIMEOUT (inline में lock नहीं हो रहा)

- bit7 PLC_COMM_LOSS (remote mode में)

# 9) UI Screens (सबको साफ दिखे ऐसा)

### A) Boot / Self-Check Screen

- “Checking ADC / Temp / Signal…”

- Result: System Ready / Self Check Failed (reason + Retry)

- Always safe: drive OFF until ready

### B) Operator Screen (Simple)

- Big: Viscosity (cP)

- Temp (°C)

- Status: Searching/Locked/Error

- Reading Health % (Signal OK %)

- Buttons (mode dependent):

  - Tabletop: Start / Stop

  - Inline: Enable / Pause (Maintenance)

- Logging: Start Log / Stop Log / Export

- Alarm Details

### C) Alarms Screen

- Active alarms list + meaning + suggested action

- Acknowledge (optional)

- Reset (Engineer/PLC command)

### D) Engineer Screen (Password)

Tabs:

- Mode & Control: Tabletop/Inline, Local/Remote/Mixed, Remote Enable, Comm-loss action

- Calibration Wizard: Add unlimited points (sample name + known cP + capture 10 sec)

- Diagnostics: ADC ok, temp ok, saturation, signal level, lock behavior

- Trends: viscosity/freq/health charts

- PLC/Network: Modbus enable, IP/port, register map view/export

## 10) Calibration (Unlimited points, easy)

- Minimum: Air + Water + 2 oils

- Best: 4–7 standards (1, 10, 50, 100, 200, 500 cP जैसे)

- Model: LUT + interpolation (safe, stable)

- Multi profiles: Profile-1 (Oil-A), Profile-2 (Oil-B)

## 11) “How to use” (1-minute SOP)

Tabletop:

- Power ON → auto open → Ready

- (जरूरत हो तो) Air calibration

- sample dip → Start → Locked → reading note/log → Stop

Inline:

- Power ON → auto open → Ready

- Remote enable ON हो तो auto run / PLC start

- Maintenance में Pause

- Alarm आए तो alarm screen + corrective action

## 12) Acceptance checklist (field readiness)

- Power fail के बाद auto-open + safe state

- Remote/Local conflict नहीं (priority rules सही)

- Overcurrent/saturation पर drive OFF + alarm

- PLC registers stable + heartbeat running

- Calibration points add करके profile switch हो सके