Predictron ViscoLogic

Expandable Software Architecture (Tabletop + Inline + PLC) | Version 2.1 | 2025-12-23

यह document ViscoLogic viscosity meter software की complete architecture देता है, ताकि development, testing, deployment, और PLC integration बिना confusion के किया जा सके।

# 1) Scope और Key Requirements

Target: Raspberry Pi based viscosity meter software जो Tabletop और Inline दोनों modes में चले, UI-friendly हो, और PLC के साथ reliably integrate हो।

## 1.1 Must-have requirements (Locked)

Auto-start on boot (power ON होते ही software खुद open हो)

Power fail recovery: restart पर safe state + last settings restore

Tabletop mode (manual Start/Stop) + Inline mode (continuous, Pause/Maintenance support)

Control Source: Local / Remote / Mixed (priority rules के साथ conflict-free)

Safety-first: drive current limit 150 mA (hardware + software), soft-start ramp, fault latch

Signal processing: resonance sweep + lock + re-lock, noise cleaning, stable reading

Reading Health % (No ML): signal/lock quality based score

Calibration: unlimited points + multiple profiles, LUT interpolation (safe default)

Temperature: PT100/MAX31865 read + optional normalization

PLC integration: Modbus TCP (data out + commands in) + heartbeat + alarms + command handshake

Logging: CSV + SQLite (recommended), retention/auto-clean

First-run Commissioning Password: नई machine पर सिर्फ पहली बार required

Operator friendly: Start button style workflow, minimum steps, clear alarms

## 1.2 Application behavior (Auto-open + manual close)

Default: systemd service boot पर app start करेगा और UI open करेगा।

अगर operator window बंद कर दे, तो system policy configurable रहेगी:

Option A (recommended for production): service auto-restart करके UI फिर से open कर दे।

Option B (lab/testing): close करने पर service stop रहे; operator manually restart कर सके।

Engineer mode में 'Shutdown/Stop Service' action (password required) provide किया जाएगा।

## 1.3 Access Levels

Operator: daily use (Start/Enable, reading, logging, export, alarms view)

Engineer (password): calibration, settings, diagnostics, commissioning reset, service control

PLC/SCADA: remote start/stop, mode select (as allowed), live values read

# 2) One-time Commissioning Lock (New Machine)

Requirement: नई machine / fresh installation पर software पहली बार run होते समय commissioning password मांगे। एक बार commissioning complete होने के बाद normal operation में password repeat नहीं होगा।

## 2.1 First-run flow

Boot → Auto-open → Self-check

अगर device_state.commissioned = false: Commissioning Lock Screen (password prompt)

Password OK → Commissioning Wizard (basic setup)

Complete Setup → commissioned=true save → normal Operator screen

## 2.2 Password storage (No plain text)

Passwords salted hash में store होंगे (bcrypt/argon2 recommended)।

Commissioning password: one-time gate (fresh device only)।

Engineer password: हमेशा required (settings/calibration access)।

## 2.3 PLC behavior during commissioning

Drive हमेशा OFF (safe)।

STATUS_WORD bit12 COMMISSIONING_REQUIRED = 1।

PLC commands ignore (या limited) until commissioned=true।

# 3) High-Level Layered Architecture

Architecture का core idea: UI और PLC दोनों एक ही orchestration layer से control करेंगे; hardware को direct touch नहीं करेंगे।

+------------------------------- UI Layer --------------------------------+

| Operator | Engineer | Calibration Wizard | Alarms | Commissioning        |

+--------------------------+----------------------------------------------+

|

v

+------------------------ Core Orchestrator (State Machine) --------------+

| BOOT/SELF_CHECK/COMMISSIONING/SAFE_IDLE/SWEEP/LOCK/RUN/FAULT             |

| Stop>Start, E-Stop highest, comm-loss action, safe defaults              |

+--------------------------+------------------------------+----------------+

|                              |

v                              v

+------------------ Measurement Pipeline ----------------+  +-------------+

| Acquisition -> DSP (Lock-in/IQ) -> Sweep/PLL -> Health |  | Protocols   |

| -> Viscosity Model (Calibration) -> Temp Compensation  |  | Modbus TCP  |

+--------------------------+------------------------------+----------------+

|

v

+------------------------- Hardware Drivers ------------------------------+

| ADS1115 (diff) | PWM Drive | PT100/MAX31865 | Safety Inputs (optional)  |

+-------------------------------------------------------------------------+

+---------------------------- Storage/Logging ----------------------------+

| SQLite (device_state, profiles, points, events) | CSV logs | retention  |

+-------------------------------------------------------------------------+

# 4) Runtime Components (Threads / Services)

Single application process के अंदर multiple threads/services चलेंगे ताकि stability और UI responsiveness बनी रहे।

Acquisition Thread: ADS1115 differential sampling + PT100 read + ring-buffer fill

Control Loop Thread (Main Engine): sweep/lock/relock, PWM control, safety checks, viscosity compute

Protocol Thread: Modbus TCP server (read registers + write commands + handshake)

UI Thread: operator/engineer screens, charts, wizard flows

## 4.1 Inter-thread communication

EventBus/Queues (thread-safe): samples_queue, commands_queue, status_pub

UI direct hardware को touch नहीं करेगी; orchestrator APIs/events के through

# 5) Core Data Models (Standardized)

## 5.1 MeasurementFrame

timestamp_ms

raw_adc_mv (or raw_adc_counts)

freq_hz

lock_state: SEARCHING / LOCKING / LOCKED / LOST

magnitude_clean

phase_deg

temp_c

viscosity_cp

reading_health_pct (0-100, No ML)

status_word, alarm_word

## 5.2 Command

source: LOCAL / PLC

cmd_type: START, STOP, PAUSE, RESET_ALARMS, SET_MODE, SET_CONTROL_SOURCE, SET_REMOTE_ENABLE

seq_id (PLC handshake)

timestamp_ms

# 6) State Machine (Industrial Reliability)

## 6.1 States

BOOTING

SELF_CHECK

COMMISSIONING_LOCK (first run only)

COMMISSIONING_WIZARD (first run only)

SAFE_IDLE (Drive OFF, ready)

SWEEPING

LOCKING

LOCKED_RUNNING

PAUSED_MAINTENANCE

FAULT_LATCHED (E-Stop/Overcurrent/Critical)

## 6.2 Key rules (Locked)

Startup में हमेशा Drive OFF until READY

Stop हमेशा Start को override करेगा

E-Stop highest priority और latched

Inline mode में comm-loss action configurable (Pause recommended)

Tabletop mode में auto-resume default OFF; Inline mode में optional ON

# 7) Measurement Pipeline (DSP + Lock + Health)

## 7.1 Acquisition

ADS1115 differential read (A0-A1) + PGA auto-range (saturation guard)

Sampling window fixed (N samples per update) ताकि DSP stable रहे

PT100 temperature read at slower rate (e.g., 2-5 Hz)

## 7.2 Resonance Sweep + Lock (Auto-tuning)

Startup sweep: f_min to f_max (configurable), best point pick by magnitude+phase

Runtime tracking: drift detect → frequency nudge → re-lock

Lost-lock: controlled retry + timeout → alarm (Inline) या operator prompt (Tabletop)

## 7.3 Reading Health % (No ML)

Inputs: lock stability, signal magnitude SNR, clipping flags, frequency drift rate

Output: 0-100 score (operator decision support)

Low health: viscosity freeze/hold-last-value option (configurable)

# 8) Safety Manager (150 mA limit)

## 8.1 Hard safety rules

Max drive current: 150 mA (absolute cap)

Air calibration low power (default 30-60 mA) + max duration limit

Soft-start ramp (step-wise PWM/current increase)

Overcurrent/overheat/fault → Drive OFF + FAULT_LATCHED

## 8.2 Software safety behaviors

ADC saturation guard: clip detect → auto PGA adjust + warning

Command safety: START allowed only if READY and not FAULT_LATCHED

Inline mode में PLC STOP हमेशा accepted (even if local disabled)

# 9) Calibration + Profiles (Unlimited points)

## 9.1 Calibration strategy

Minimum recommended: Air + Water (1 cP) + 2-3 standard oils

Best practice: 4-7 points (1, 10, 50, 100, 200, 500 cP etc.)

Multiple profiles supported (Oil-A, Oil-B, Product-1...)

## 9.2 Default model

LUT + interpolation (robust, avoids wild extrapolation)

Optional: polynomial fit (engineer-only, with validation guards)

## 9.3 Calibration Wizard flow

Profile select/create

Add Point: sample name + known viscosity (cP)

Capture window (10 sec) → stable value auto pick

Save point → LUT update

Validate: residual error + outlier warnings

# 10) Control Arbitration (Local/Remote/Mixed)

Problem: Local और PLC दोनों start/stop दे सकते हैं, conflict avoid करना जरूरी है।

## 10.1 ControlSource definition

Local: सिर्फ local UI से run control, PLC only monitoring

Remote: PLC commands active, UI run control disabled (monitoring only)

Mixed: दोनों allowed, लेकिन strict priority rules

## 10.2 Priority rules (Locked)

STOP always wins (local stop या PLC stop में से जो भी आए)

FAULT_LATCHED में START ignore

Mixed mode में 'Last command wins' नहीं; rules deterministic होंगी:

Remote start valid only if RemoteEnable=1

Local start valid only if LocalEnable=1 (engineer configurable)

# 11) PLC Integration (Modbus TCP) - Register Map

Default: Raspberry Pi = Modbus TCP Server, PLC = Modbus Client. Address offset (0-based vs 1-based) PLC brand के हिसाब से adjust होगा।

## 11.1 Holding Registers - PLC Reads (4xxxx)

| Human Addr | Name | Type | Scaling | Notes |
| --- | --- | --- | --- | --- |
| 40001-40002 | VISC_CP_x100 | int32 | cP*100 | 45.23 cP => 4523 |
| 40003 | TEMP_C_x100 | int16 | C*100 | 42.15C => 4215 |
| 40004 | FREQ_HZ_x100 | int16 | Hz*100 | 178.95Hz => 17895 |
| 40005 | READING_HEALTH | uint16 | 0-100 | No ML quality score |
| 40006 | STATUS_WORD | uint16 | bitfield | state + flags |
| 40007 | ALARM_WORD | uint16 | bitfield | active alarms |
| 40008 | HEARTBEAT | uint16 | +1/sec | rollover OK |
| 40009 | LAST_CMD_SEQ | uint16 | - | last processed cmd seq |
| 40010 | MAPPING_VERSION | uint16 | - | e.g., 21 for v2.1 |

## 11.2 Holding Registers - PLC Writes (41xxx)

| Human Addr | Name | Type | Meaning |
| --- | --- | --- | --- |
| 41001 | MODE_SELECT | uint16 | 0=Tabletop, 1=Inline |
| 41002 | CONTROL_SOURCE | uint16 | 0=Local, 1=Remote, 2=Mixed |
| 41003 | REMOTE_ENABLE | uint16 | 0=Disable remote run, 1=Enable |
| 41004 | CMD_WORD | uint16 | bit commands (start/stop/pause/reset) |
| 41005 | CMD_SEQ | uint16 | PLC increments on new command |
| 41006 | LOG_ENABLE | uint16 | 0=Off, 1=On |
| 41007 | PROFILE_ID | uint16 | select calibration profile |

## 11.3 CMD_WORD bits (PLC to Meter)

bit0 START_REQUEST

bit1 STOP_REQUEST

bit2 PAUSE_MAINTENANCE

bit3 RESET_ALARMS

bit4 CLEAR_ESTOP_LATCH (only if safe)

## 11.4 STATUS_WORD bits (Meter to PLC)

bit0 READY

bit1 SELF_CHECK_FAIL

bit2 SEARCHING

bit3 LOCKED

bit4 LOST_LOCK

bit5 RUNNING (drive active)

bit6 PAUSED / MAINTENANCE

bit7 REMOTE_ACTIVE

bit8 ALARM_ACTIVE

bit9 ESTOP_LATCHED

bit10 ADC_SATURATION

bit11 TEMP_SENSOR_FAULT

bit12 COMMISSIONING_REQUIRED

## 11.5 Command Handshake

PLC writes CMD_WORD and CMD_SEQ

Meter processes only when CMD_SEQ changes

Meter updates LAST_CMD_SEQ (40009) and status flags accordingly

# 12) UI Screens (Operator friendly + Engineer safe)

## 12.1 Startup / Self-check

Screen: Checking ADC / Temp / Signal

Result: System Ready OR Self Check Failed (reason + Retry)

Always safe: drive OFF until READY

## 12.2 Commissioning screens (First run only)

Commissioning Lock Screen: password prompt

Commissioning Wizard: mode, control source, remote enable, PLC enable, safety defaults, logging defaults

## 12.3 Operator screen (Simple)

Big number: Viscosity (cP)

Temperature (C)

Status: Searching/Locked/Error

Reading Health %

Buttons: Start/Stop (Tabletop), Enable/Pause (Inline), Log ON/OFF, Export, Alarm Details

## 12.4 Engineer screen (Password)

Calibration Wizard + profile management

Diagnostics: ADC/PT100/lock/saturation flags

Settings: sweep range, averaging, comm-loss action, auto-resume

PLC/Network view (Modbus enable, port, mapping version)

Re-Commission device (commissioned=false) - restricted action

# 13) Storage, Logging, Retention

## 13.1 SQLite (recommended tables)

device_state: commissioned, commissioned_at, last_mode, last_control_source, remote_enable, last_profile_id

cal_profiles: id, name, created_at

cal_points: profile_id, known_cp, measured_feature, temp_c, timestamp

events: alarms, state transitions, commands (lightweight audit)

## 13.2 CSV logs

Per-run or per-day CSV (configurable)

Columns: timestamp, viscosity_cp, temp_c, freq_hz, health_pct, status_word, alarm_word

Retention policy: keep last N days, auto-delete older

# 14) Config (settings.yaml) + Validation

Mode defaults: tabletop/inline

ControlSource defaults: local/remote/mixed

CommLossAction: pause / keep_running

Drive limits: max_current_ma, ramp_ms, air_cal_current_ma

Sweep: f_min, f_max, step_hz, fine_step_hz

Logging: enable, path, retention_days

Modbus: port, unit_id, mapping_version

Config schema validation (schema.json) ताकि गलत values crash न कराएं

# 15) Suggested Folder Structure (Production-ready)

viscologic/

app.py

core/

orchestrator.py

state_machine.py

event_bus.py

config_manager.py

safety_manager.py

diagnostics.py

security/

commissioning_manager.py

auth_manager.py

drivers/

adc_ads1115.py

drive_pwm.py

temp_max31865.py

gpio_helpers.py

dsp/

lockin_iq.py

sweep_tracker.py

filters.py

health_score.py

model/

calibration_store.py

calibration_lut.py

viscosity_compute.py

temp_compensation.py

protocols/

modbus_server.py

register_map.py

storage/

sqlite_store.py

csv_logger.py

retention.py

ui/

main_window.py

operator_screen.py

engineer_screen.py

commissioning_lock_screen.py

commissioning_wizard.py

calibration_wizard.py

alarms_screen.py

trends_view.py

tests/

test_lockin.py

test_sweep.py

test_register_map.py

test_state_machine.py

deploy/

viscologic.service

kiosk_setup.md

config/

settings.yaml

schema.json

docs/

SOP_operator.md

SOP_engineer.md

PLC_register_map.md

Calibration_guide.md

# 16) Deployment (Auto-open + Auto-restart)

systemd service: boot पर auto-start, crash पर auto-restart

optional kiosk mode: fullscreen UI, no desktop distractions

safe shutdown handling: exit पर drive OFF

# 17) Testing & Simulation

Unit tests: DSP, sweep/lock, state machine rules, register map packing

Simulated drivers: fake ADC waveform generator (air/water/oil)

Soak test: 24-72 hour run, comm-loss scenarios, power-cycle tests

PLC integration test: modbus client script + command handshake verification

# 18) Implementation Roadmap (Recommended build order)

Boot auto-start + self-check + safe idle + commissioning lock

ADS1115 differential read + PGA auto-range + PT100 read

Sweep + lock + relock + stable viscosity pipeline + health score

Operator UI + alarms + logging + export

Calibration wizard + profiles + LUT model

Modbus TCP server + register map + handshake + arbitration rules

Long-run testing + packaging (kiosk/service) + documentation

# 19) One-line promise

Operator को सिर्फ Start/Enable दबाना है, बाकी software खुद signal clean करके, resonance lock करके, stable viscosity number दिखाएगा, data save करेगा और PLC को live output देगा।